# OpenELIS ASTM Analyzer Mapping Feature

## Comprehensive Specification Document

Version: 1.0

Date: November 12, 2025

Status: Draft for Review

Technology: Java Spring Framework, Carbon React

Protocol: ASTM LIS2-A2 (TCP/IP)

# Table of Contents

1. Executive Summary

2. Functional Requirements

3. Business Rules

4. User Stories

5. UI Design Guidance (Carbon Design System)

6. API Endpoints

7. Validation Rules

8. Security & Permissions

9. Testing Requirements

10. Implementation Phases

11. Success Metrics

12. Risks & Mitigations

13. Future Enhancements

14. Appendices

# 1. Executive Summary

This specification defines the analyzer field mapping feature for OpenELIS, enabling dynamic configuration of data mappings between clinical laboratory analyzers using ASTM LIS2-A2 protocol and the OpenELIS data model. This feature extends the existing ASTM communication layer with a user-friendly configuration interface built with Carbon Design System.

## Key Objectives

Enable dynamic field mapping between analyzers and OpenELIS without code changes

Support multiple analyzer types with per-analyzer or shared configurations

Provide intuitive UI conforming to Carbon Design System standards

Ensure data integrity through comprehensive validation rules

Facilitate troubleshooting with comprehensive error logging and dashboard

## Scope

This specification covers the mapping configuration interface only. The existing ASTM LIS2-A2 TCP/IP communication layer, order request handling, and result receiving mechanisms are already implemented and are not in scope for this feature.

# 2. Functional Requirements

## 2.1 Analyzer Management

### FR-1: Analyzer Registration

The system shall allow users to register new analyzers with the following attributes:

Analyzer name (unique, required, 1-100 characters)

Analyzer type/model (required, dropdown with "Other" option)

IP address (required, IPv4 format validation)

Port number (required, 1-65535 range)

Protocol version (default: ASTM LIS2-A2)

Active status (boolean, default: true)

Test unit assignment (required, multi-select, minimum 1)

Created date/time (auto-generated)

Last modified date/time (auto-generated)

Created by user (auto-generated)

### FR-2: Analyzer List View

The system shall display all analyzers in a searchable, filterable data table with:

Columns: Name, Type, IP:Port, Status, Test Units, Last Modified

Default sort: Last Modified (descending)

Pagination: 25, 50, 100 items per page

Inline status toggle (active/inactive)

### FR-3: Analyzer Search and Filter

The system shall provide:

Search: Free text search across Name, Type, IP fields (debounced, 300ms)

Filters:

Status: All, Active, Inactive

Test Unit: Multi-select dropdown

Analyzer Type: Multi-select dropdown

Filter pill display with clear functionality

Persistent filter state during session

### FR-4: Analyzer Actions

The system shall support:

Add Analyzer: Modal dialog for new analyzer creation

Edit Analyzer: Modal dialog for existing analyzer modification

Delete Analyzer: Soft delete with confirmation (only if no active mappings)

Test Connection: Verify TCP connectivity and ASTM handshake

Configure Mappings: Navigate to field mapping interface

## 2.2 Field Discovery

### FR-5: Analyzer Query

The system shall query the analyzer to retrieve available data fields:

Trigger: Manual "Query Analyzer" button in mapping interface

Process: Send ASTM query message, parse response records

Extract: All field identifiers from ASTM message segments

Display: Available fields in source panel with field type indicators

### FR-6: Field Type Detection

The system shall detect or allow specification of field types:

Numeric (with unit of measure)

Qualitative (text/interpretation)

Control Test (flag)

Melting Point (temperature)

Date/Time

Text (general)

Custom (extensible)

## 2.3 Field Mapping Configuration

### FR-7: Mapping Interface

The system shall provide a dual-panel mapping interface with:

Left Panel (Source): Searchable list, field type indicators, ASTM references

Right Panel (Target): OpenELIS fields categorized by entity, searchable, type-filtered

### FR-8: Mapping Creation

Drag-and-drop or click-to-map with visual connections and type validation

### FR-9: Unit of Measure Mapping

For numeric fields: display analyzer unit, select OpenELIS unit, flag mismatches

### FR-10: Qualitative Result Mapping

Map analyzer values to OpenELIS codes, support many-to-one, default handling

### FR-11: Copy Mappings

Copy all field mappings from another analyzer with search and confirmation

### FR-12: Custom Field Types

Administrators can add new field types with validation rules

### FR-13: Dynamic Target Fields

Add new OpenELIS target fields directly from mapping interface

# 3. Business Rules

## BR-1: Analyzer Uniqueness

Analyzer name must be unique across the system

IP:Port combination generates warning if duplicate

## BR-2: Active Status

Only active analyzers receive orders

Inactive analyzers retain configuration

Status changes logged in audit trail

## BR-3: Test Unit Assignment

Minimum one test unit required

Test units control visibility in filtered views

## BR-4: Mapping Validation

Required mappings: Sample ID, Test Code, Result Value

Missing required mappings prevent activation

## BR-5: Field Type Compatibility

Error: text to numeric (blocked)

Warning: numeric to text (allowed with confirmation)

Error: missing unit for numeric fields (blocked)

## BR-6: Mapping Changes

Active analyzer changes require confirmation

Changes apply to new results only

Complete audit log maintained

## BR-7: Delete Protection

Analyzers with recent results cannot be hard deleted

Soft delete only (90-day window)

## BR-8: Connection Testing

Validates TCP handshake only

Timeout: 30 seconds

Displays latency and timestamp

## BR-9: Query Limitations

Query timeout: 5 minutes

Maximum 500 fields per query

Rate limit: 1 query per minute per analyzer

## BR-10: Copy Mappings Validation

Source must have active mappings

Overwrites existing mappings with confirmation

Type incompatibilities generate warnings

# 4. User Stories

## Epic 1: Analyzer Setup

### US-1.1: Register New Analyzer

As a laboratory administrator
I want to register a new analyzer in the system
So that I can receive results from that instrument

Acceptance Criteria:

Given I'm on the Analyzers page

When I click "Add Analyzer"

Then I see a modal with connection and configuration fields

And I can save the analyzer after validation

### US-1.2: Test Analyzer Connection

As a laboratory administrator
I want to test connectivity to an analyzer
So that I can verify the configuration before going live

Acceptance Criteria:

Given I have entered analyzer connection details

When I click "Test Connection"

Then the system attempts TCP connection

And displays success/failure with latency

### US-1.3: Manage Analyzer Status

As a laboratory supervisor
I want to activate or deactivate analyzers
So that I can control which instruments receive orders

Acceptance Criteria:

Given I'm viewing the analyzer list

When I toggle the status switch

Then the analyzer status changes immediately

And change is logged in audit trail

## Epic 2: Field Mapping

### US-2.1: Query Analyzer Fields

As a laboratory administrator
I want to query an analyzer for its available data fields
So that I know what data can be mapped

Acceptance Criteria:

When I click "Query Analyzer"

Then the system sends ASTM query

And displays all available fields

### US-2.2: Map Analyzer Fields

As a laboratory administrator
I want to map analyzer fields to OpenELIS data fields
So that results are correctly stored

Acceptance Criteria:

When I drag a source field to target

Then a visual connection is created

And mapping is validated for compatibility

### US-2.3: Map Numeric Results

As a laboratory administrator
I want to specify units of measure for numeric results
So that values are correctly interpreted

Acceptance Criteria:

When I create numeric mapping

Then I see unit dropdown

And unit mapping is saved

### US-2.4: Map Qualitative Results

As a laboratory administrator
I want to map qualitative interpretations
So that text results are standardized

Acceptance Criteria:

When I create qualitative mapping

Then I can map each value to OpenELIS code

And specify default for unmapped values

### US-2.5: Copy Mappings

As a laboratory administrator
I want to copy field mappings from similar analyzer
So that I don't reconfigure identical instruments

Acceptance Criteria:

When I click "Copy from existing analyzer"

Then I see searchable list of analyzers

And can confirm to copy all mappings

## Epic 3: Analyzer Management

User stories for searching, filtering, viewing details, and editing analyzer configurations.

## Epic 4: System Administration

User stories for adding custom field types and viewing the error dashboard.

# 5. UI Design Guidance (Carbon Design System)

All user interfaces must conform to IBM Carbon Design System standards. This ensures consistency, accessibility, and professional appearance.

## 5.1 Analyzers List Page

Layout Structure:

Page header with "Analyzers" title and "Add Analyzer" primary action button

Search bar with filter dropdowns (Status, Test Unit, Analyzer Type)

Active filter pills below search bar

Data table with sortable columns and inline actions

Pagination controls at bottom

Carbon Components:

<Header> - Page title and primary actions

<Search> - Text search (300ms debounce)

<MultiSelect> - Filter dropdowns

<Tag> - Filter pills (dismissible)

<DataTable> - Analyzer listing

<Toggle> - Status control

<OverflowMenu> - Row actions

<Pagination> - Navigation controls

## 5.2 Add/Edit Analyzer Modal

Modal Configuration:

Type: <ComposedModal> medium size (640px)

Scrollable content with fixed action buttons

Form fields: <TextInput>, <Dropdown>, <NumberInput>, <MultiSelect>

Connection test button (ghost) in fieldset

Action buttons: Cancel (secondary), Save (primary)

## 5.3 Field Mapping Interface

Two-Column Layout:

Equal width panels (50/50) using CSS Grid. Stack vertically on mobile (<1024px).

Source Panel (Left):

<Accordion> for field categories

Checkbox selection with field type <Tag>

ASTM segment reference (light gray)

Status bar: "X mapped • Y available"

Target Panel (Right):

<Accordion> for entity categories

Checkboxes (checked when mapped)

Inline <Dropdown> for unit selection

Visual connection lines between mapped fields

Field Type Colors:

| Type | Color | Token |
| --- | --- | --- |
| Numeric | Blue | $blue-60 |
| Qualitative | Purple | $purple-60 |
| Control | Green | $green-60 |
| Melting Point | Teal | $teal-60 |
| Date/Time | Cyan | $cyan-60 |
| Text | Gray | $gray-60 |

## 5.4 Carbon Design Principles

Color System:

Use Carbon color tokens only. Never hard-code hex values.

Typography:

Page titles: $heading-04 (32px)

Section headers: $heading-03 (20px)

Body text: $body-01 (14px)

Helper text: $label-01 (12px)

Spacing:

Use Carbon spacing tokens: $spacing-05 (16px) standard, $spacing-07 (32px) sections

Accessibility (WCAG 2.1 AA):

Keyboard navigation for all interactive elements

Screen reader labels for icon buttons

Color contrast ratio ≥ 4.5:1

Focus indicators on all controls

# 6. API Endpoints

All API endpoints follow RESTful conventions and return JSON responses. Authentication required using OpenELIS security framework.

## Analyzer Management

GET /api/analyzers

Retrieve paginated list with filters

Query: search, status, testUnit[], page, pageSize

Authorization: LAB_USER

POST /api/analyzers

Create new analyzer

Body: name, analyzerType, ipAddress, port, testUnitIds

Authorization: LAB_ADMIN

PUT /api/analyzers/{id}

Update existing analyzer

Body: Same as POST, all fields optional

Authorization: LAB_ADMIN

DELETE /api/analyzers/{id}

Soft delete analyzer

Validates no recent results

Authorization: LAB_ADMIN

POST /api/analyzers/{id}/test-connection

Test TCP connection

Returns: success, latencyMs, timestamp

Authorization: LAB_ADMIN

## Field Discovery & Mapping

POST /api/analyzers/{id}/query-fields

Query analyzer for available fields

5-minute timeout, returns array of field definitions

Authorization: LAB_ADMIN

GET /api/analyzers/{id}/mappings

Retrieve all mappings

Returns array of mapping objects

Authorization: LAB_USER

POST /api/analyzers/{id}/mappings

Create mappings (batch)

Body: Array of mapping definitions

Authorization: LAB_ADMIN

POST /api/analyzers/{id}/mappings/copy

Copy from another analyzer

Body: sourceAnalyzerId

Authorization: LAB_ADMIN

POST /api/analyzers/{id}/mappings/{mid}/qualitative-values

Map qualitative values

Body: Array of value mappings

Authorization: LAB_ADMIN

## Error Dashboard

GET /api/analyzer-errors

Retrieve errors with filters

Query: analyzerId, errorType, startDate, endDate, page

Authorization: LAB_SUPERVISOR

POST /api/analyzer-errors/{id}/acknowledge

Acknowledge error

Returns 204 No Content

Authorization: LAB_SUPERVISOR

POST /api/analyzer-errors/acknowledge-all

Acknowledge multiple

Body: Optional filters

Authorization: LAB_SUPERVISOR

# 7. Validation Rules

## Analyzer Validation

| Field | Rule | Error Message |
| --- | --- | --- |
| Name | 1-100 chars, unique | Name required and must be unique |
| Type | Required | Type is required |
| IP Address | Valid IPv4 | Valid IP required (e.g., 192.168.1.10) |
| Port | 1-65535 | Port must be 1-65535 |
| Test Units | Min 1 | At least one test unit required |

## Field Mapping Validation

| Field | Rule | Message |
| --- | --- | --- |
| Source | Required, unique | Source field required and unique |
| Target | Required | Target field required |
| Unit | Required for numeric | Unit of measure required |
| Required Maps | Sample ID, Test Code, Result | Missing required mappings |

# 8. Security & Permissions

## Role-Based Access Control

| Role | Permissions |
| --- | --- |
| RESULTS | View analyzers, mappings, errors |
| SYSTEM_ADMIN | All permissions + Create field types |

## Audit Trail

All administrative actions logged with: action type, entity, user, timestamp, IP, before/after values.

# 9. Testing Requirements

# 10. Testing Requirements

## Unit Tests

All service methods

Validation logic

Type compatibility

Coverage: 80%

## Integration Tests

API endpoints

Database transactions

ASTM integration

Coverage: 70%

## UI Tests

Form validation

Search/filter

Mapping interface

Modal interactions

## UAT

Complete workflows

Real analyzer testing

Laboratory staff validation

## Performance

1000 analyzers load test

Query timeout handling

Concurrent users

# 13. Future Enhancements (Out of Scope)

## Automated Field Matching

ML-based mapping suggestions from historical data

## Bi-directional Sync

Real-time status updates from analyzers

## Advanced Transformations

Complex transformation rules with formulas

## Multi-Protocol Support

HL7, FHIR in addition to ASTM

## Analyzer Groups

Manage identical analyzers with shared mappings

## Visual Data Flow

Graphical representation of data flow

## Import/Export

Backup and share configurations as JSON/XML

## Scheduled Queries

Automatic periodic capability queries

## Analytics Dashboard

Utilization and performance metrics

# 14. Appendices

## Appendix A: ASTM LIS2-A2 Field Reference

Common ASTM segments and fields for mapping reference.

Patient (P) Segment:

P|1: Sequence Number

P|3: Patient ID

P|6: Patient Name

P|8: Date of Birth

P|9: Patient Sex

Order (O) Segment:

O|1: Sequence Number

O|2: Sample ID / Accession Number

O|4: Universal Test ID

O|6: Ordered Date/Time

O|15: Sample Type

Result (R) Segment:

R|1: Sequence Number

R|2: Universal Test ID

R|3: Data/Measurement Value

R|4: Units

R|6: Result Abnormal Flag

R|7: Result Status

Comment (C) Segment:

C|1: Sequence Number

C|2: Comment Source

C|3: Comment Text

## Appendix B: Carbon Design System Resources

Component Library: https://carbondesignsystem.com/components/overview/

Icons: https://carbondesignsystem.com/guidelines/icons/library/

Design Patterns: https://carbondesignsystem.com/patterns/overview/

React Package: @carbon/react npm package

## Appendix C: Glossary

ASTM: American Society for Testing and Materials

LIS2-A2: Laboratory Instrument Standard 2, version A2

Analyzer: Clinical laboratory instrument for automated testing

Field Mapping: Association between analyzer and OpenELIS fields

Test Unit: Organizational laboratory unit (e.g., Hematology)

Qualitative Result: Non-numeric result (e.g., Positive/Negative)

Unit of Measure: Standard measurement unit (e.g., mg/dL)

Carbon Design System: IBM open-source design system

Soft Delete: Marking record deleted without physical removal

## Appendix D: Common Units of Measure

Concentration:

mg/dL

g/dL

mmol/L

μmol/L

ng/mL

IU/L

Cell Counts:

10^3/μL

10^6/μL

10^9/L

cells/μL

%

Temperature:

°C

°F

Time:

seconds

minutes

hours

Volume:

mL

μL

L